|  |  |  |  |  |  |  | The maximum mass fraction of hazardous elements designated in the certificate of analysis: Element; Mass fraction. Cd; ≤ 0.002%. Be; ≤ 0.002%. Pb; ≤ 0.002%. Ni; ≤ 0.01%. See CoA, Doc-ID LLxxxx |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Bio-compatibility | 5.3 | Biocompatibility | Y/A | Assessment acc. particular standards for biocompatibility |  | ISO zzzzz-1 ISO a2222 | The assessment is handled as a separate detached document. See “6 - test Assessment – ProductXYZ”, Doc-ID LLxxxx |  |  |  |  |  | PASS |
| Mechanical properties | 5.4.1 | General: test | Y/A | After - test - test2-cooling, - test3, - test4 |  | ISO aaaa ISO bbbb-1 | The acceptance criteria for the mechanical properties have been met for using the material as intended |  |  |  |  |  | PASS |
|  | 5.4.2 | Proof stress of 0,2 % test (Rp0.2) | Y/A | ≥ 360 MPa (Type 4) |  | ISO zzz ISO yyy-1 | 1st Batch "zz_1_zz" (n=6 samples) Rp0.2 mean =520.0 MPa, stdev = 10.3 MPa |  |  |  |  |  | PASS |
|  |  |  |  | Pass | Fail |  | 111 | 222 | 333 | 444 | 555 | 888 |  |
|  |  |  |  | 1st occasion (first lot of 6) |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  | PASS |
|  | 5.4.3 | test after test, (test) | Y/A | ≥ 2% (Type 4) |  | ISO eee ISO sss-1 |  |  |  |  |  |  | PASS |
|  |  |  |  |  |  |  | 8.0 | 8.9 | 12.8 | 6.6 | 6.9 | 8.2 |  |
|  | 5.5 | test | N/A | N/A for Type 0-4 |  | ISO yyy ISO uuu-1 method |  |  |  |  |  |  | PASS |
|  |  |  |  |  |  |  | 229 | 274 | 261 | 230 |  |  |  |

# History

Record/document creation: Please delete the template change history.

Reason for the preparation / revision: test info

Replaces the following documents: V1.0